2010  PROTOCOL:

- circle one or more resources:

HUP5 1.5T      HUP6 3T        DEVON 3T           Glenolden 3T         HUP7 7T

- circle one or more pulse sequences:

(Note that red type = experimental pulse sequence and blue type = standard pulse sequence.).

# DESCRIPTION	  SEQUENCE ACQUISITIONS

- SAGGITAL

SCOUT

Plane  Saggital	Coil   Siemens ARRAY or HEAD  Psd File Name gre

Tr       20ms	Te    5ms		Fov   28cm

Thickness   8mm	Skip  0	Soft

Matrix   192x144	Nex    1	Center Freq. 1H

2. 3D-HIGH-RES.

MPRAGE

Plane  AXIAL	Coil   Siemens ARRAY or HEAD  Psd File Name mpr

Tr       1620ms	Te    3.87ms		Fov   25cm

Thickness   1mm	Skip  0	Soft

Matrix   256X256	Nex    1	 	Center Freq. 1H

3. FLAIR-T2

Plane  AXIAL		Coil   Siemens ARRAY or HEAD  	Psd File Name tse

Tr       4000ms		Te    85ms	Fov   24cm

Thickness   5mm	Skip  1.5mm	 Soft

Matrix   256X256	Nex    1	 Center Freq. 1H

4. BOLD fMRI

Plane  AXIAL		Coil   Siemens ARRAY or HEAD  	Psd Filename ep2d

Tr       3000ms		Te    30ms	Fov   20cm

Thickness   3mm	Skip  0	 Soft

Matrix  64x64	Nex    1	 Center Freq. 1H

5. PASL

Plane  AXIAL	Coil   Siemens ARRAY or HEAD    Psd File Name		                                                  ep2D_fairest_UI_ipat2

Tr       3000ms	Te    17ms		Fov   24cm

Thickness   6mm	Skip  1.5mm	Soft

Matrix   64X64	Nex    8	Center Freq.

1Hpost IR delay 1400-1600ms

Post inf sat delay 600ms

1R slab thickness 100-150mm

Ns1R slab thickness 300-350mm

6. 	    6. pseudo-CASL

Plane  AXIAL  Coil   Siemens ARRAY or HEAD  Psd File Name 				ep2d_pcasl_ul_phc

Tr       4000ms	Te    17ms	         Fov   24cm

Thickness   6mm	Skip  1.5mm                           Soft

Matrix   64X64	Nex    80	 Center Freq. 1H

Label offset 80-90mm	num RF blocks 50-82

Post label delay 1000-1500ms	RF gap 360ms

Mean Gzx10 6mt/m

Phi adjust  100%

7. CASL

Plane  AXIAL	Coil  Siemens ARRAY or HEAD    Psd File Name ep2D_casl_UI_ipat2

Tr       3000ms		Te    17ms			  Fov 24cm

Thickness   6mm	Skip  1.5mm		Soft

Matrix   64X64	Nex    80	Ctr. Freq. 1H

Label offset 80mm		                               Num RF Blocks 20

Post label delay 1000-1500ms	                               RF gap 500ms

8. DTI

Plane  AXIAL	Coil   Siemens ARRAY or HEAD  	Psd File Name DTI

Tr       5600ms	Te    87ms		Fov   24cm

Thickness   3mm	Skip  interleave	Soft ep2d_diff

Matrix   128x128     Nex    6	                                   Center Freq. water